Testing 1D Motion Platform for 6D-IEC framework experiment

Royal North Shore Hospital

03/07/2024

This report shows the data acquired by D415 RealSense camera and by the 6DoF robot feedback during the experiment conducted by Alicja Kaczynska at Royal North Shore Hospital (03/07/2024). The motion traces that have been tested can be found on ShareDrive: Y:\PRJ-RPL\2RESEARCH\2_ProjectData\IEC\Motion_traces\Traces_Jun2024_expt

The three Simulated 1D traces for rotation (AP, SI, LR) are the same input file.

There are no statistics for 6D motion traces because the accuracy of the 6DoF platform is already well known and has already been tested in a clinical environment. Furthermore, the two motions have not been recorded by the same device so the two measurements are independent.

The “XX_resp_XX.txt” files are the 1DoF motion traces.

Contents

1 DoF Platform overall accuracy	2

1.Mean_Motion_resp_shifted_gradual_start	3

1.Mean_Motion_gradual_start	4

2.Mean_Motion2_resp_shifted_gradual_start	5

2.Mean_Motion_gradual_start	6

3.Mean_Motion3_resp_shifted_gradual_start	7

3.Mean_Motion_gradual_start	8

4.Mean_Motion4_resp_shifted_gradual_start	9

4.Mean_Motion_gradual_start	10

Simulated_1D_trace_for_rotation (AP)	11

Simulated_6D_trace_for_rotation (AP)	12

Simulated_1D_trace_for_rotation (SI)	13

Simulated_6D_trace_for_rotation (SI)	14

Simulated_1D_trace_for_rotation (LR)	15

Simulated_6D_trace_for_rotation (LR)	16

# 1 DoF Platform overall accuracy

| Trace | 1DoF Motion traces |
| --- | --- |
| Mean (mm) | -0.2 |
| Std (mm) | 0.7 |
| 1st percentile (mm) | -1.7 |
| 99th percentile (mm) | 1.4 |

Figure 0: Experimental setup

# 1.Mean_Motion_resp_shifted_gradual_start

| Trace | 1.Mean_Motion_resp_shifted_gradual_start |
| --- | --- |
| Mean (mm) | -0.2 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.2 |
| 99th percentile (mm) | 1.8 |

Figure 1.1: 1.Mean_Motion resp motion trace

# 1.Mean_Motion_gradual_start

| Figure 1.2: 1.Mean_Motion motion trace |  |
| --- | --- |

# 2.Mean_Motion2_resp_shifted_gradual_start

| Trace | 2.Mean_Motion2_resp_shifted_gradual_start |
| --- | --- |
| Mean (mm) | -0.3 |
| Std (mm) | 0.8 |
| 1st percentile (mm) | -2.2 |
| 99th percentile (mm) | 1.7 |

Figure 2.1: 2.Mean_Motion2 resp motion trace

# 2.Mean_Motion_gradual_start

| Figure 2.2: 2.Mean_Motion motion trace |  |
| --- | --- |

# 3.Mean_Motion3_resp_shifted_gradual_start

| Trace | 3.Mean_Motion3_resp_shifted_gradual_start |
| --- | --- |
| Mean (mm) | -0.1 |
| Std (mm) | 0.7 |
| 1st percentile (mm) | -1.7 |
| 99th percentile (mm) | 1.2 |

Figure 3.1: 3.Mean_Motion3 resp motion trace

# 3.Mean_Motion_gradual_start

| Figure 3.2: 3.Mean_Motion motion trace |  |
| --- | --- |

# 4.Mean_Motion4_resp_shifted_gradual_start

| Trace | 4.Mean_Motion4_resp_shifted_gradual_start |
| --- | --- |
| Mean (mm) | -0.0 |
| Std (mm) | 0.7 |
| 1st percentile (mm) | -1.7 |
| 99th percentile (mm) | 1.4 |

Figure 4.1: 4.Mean_Motion4 resp motion trace

# 4.Mean_Motion_gradual_start

| Figure 4.2: 4.Mean_Motion motion trace |  |
| --- | --- |

# Simulated_1D_trace_for_rotation (AP)

| Trace | Simulated 1D trace for rotation (AP) |
| --- | --- |
| Mean (mm) | -0.6 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.8 |
| 99th percentile (mm) | 0.8 |

Figure 5.1: Simulated 1D rotation motion trace (AP)

# Simulated_6D_trace_for_rotation (AP)

| Figure 5.2: Simulated 6D rotation motion trace (AP) |  |
| --- | --- |

# Simulated_1D_trace_for_rotation (SI)

| Trace | Simulated 1D trace for rotation (SI) |
| --- | --- |
| Mean (mm) | -0.1 |
| Std (mm) | 0.7 |
| 1st percentile (mm) | -1.5 |
| 99th percentile (mm) | 1.5 |

Figure 6.1: Simulated 1D rotation motion trace (SI)

# Simulated_6D_trace_for_rotation (SI)

| Figure 6.2: Simulated 6D rotation motion trace (SI) |  |
| --- | --- |

# Simulated_1D_trace_for_rotation (LR)

| Trace | Simulated 1D trace for rotation (LR) |
| --- | --- |
| Mean (mm) | -0.1 |
| Std (mm) | 0.7 |
| 1st percentile (mm) | -1.7 |
| 99th percentile (mm) | 1.5 |

Figure 7: Simulated 1D rotation motion trace (LR)

# Simulated_6D_trace_for_rotation (LR)

| Figure 6.2: Simulated 6D rotation motion trace (SI) |  |
| --- | --- |