Combined system 1D-UR16

The aim of this report is to caracterize the accuracy of the UR16 6DoF platform and identify the time synchronization of the combined system.The motion traces that have been tested can be found on ShareDrive: Y:\PRJ-RPL\2RESEARCH\2_ProjectData\IEC\Motion_traces\Traces_Jun2024_expt

The “XX_resp_XX.txt” files are the 1DoF motion traces.

Contents

Combined time synchronization	2

UR16 performance	2

1.Mean_Motion1 (test1)	3

1.Mean_Motion1 (test2)	4

1.Mean_Motion1 (test3)	5

2.Mean_Motion (test1)	6

2.Mean_Motion (test2)	7

2.Mean_Motion2 (test3)	8

3.Mean_Motion_1 (test1)	9

3.Mean_Motion_1 (test2)	10

3.Mean_Motion_1 (test3)	11

4.Mean_Motion_2 (test1)	12

4.Mean_Motion_2 (test2)	13

4.Mean_Motion_2 (test3)	14

1.Mean_Motion1 (test 1100s)	15

# Combined time synchronization

| (s) | Combined system latency |
| --- | --- |
| Mean | 0.14 |
| Std | 0.06 |
| 1st percentile | 0.02 |
| 99th percentile | 0.28 |

# UR16 performance

| (mm) | UR 16 |
| --- | --- |
| Mean | -0.0 |
| Std | 0.6 |
| 1st percentile | -1.4 |
| 99th percentile | 1.4 |

# 1.Mean_Motion1 (test1)

| Trace | 1.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | 0.0 |
| Std (mm) | 0.7 |
| 1st percentile (mm) | -1.7 |
| 99th percentile (mm) | 2.0 |
| Trace | 1.Mean_Motion_6D |
| Mean (mm) | -0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.02 |

Figure 1.1: 1.Mean_Motion motion trace test 1

# 1.Mean_Motion1 (test2)

| Trace | 1.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | 0.0 |
| Std (mm) | 0.8 |
| 1st percentile (mm) | -1.9 |
| 99th percentile (mm) | 1.8 |
| Trace | 1.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.7 |
| 1st percentile (mm) | -1.5 |
| 99th percentile (mm) | 1.5 |
| Latency (s) | 0.13 |

Figure 1.2: 1.Mean_Motion motion trace test 2

# 1.Mean_Motion1 (test3)

| Trace | 1.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | -0.0 |
| Std (mm) | 0.8 |
| 1st percentile (mm) | -1.8 |
| 99th percentile (mm) | 2.2 |
| Trace | 1.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.14 |

Figure 1.3: 1.Mean_Motion1 motion trace test 3

# 2.Mean_Motion (test1)

| Trace | 2.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | 0.0 |
| Std (mm) | 1.0 |
| 1st percentile (mm) | -2.4 |
| 99th percentile (mm) | 2.4 |
| Trace | 2.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.3 |
| 99th percentile (mm) | 1.3 |
| Latency (s) | 0.10 |

Figure 2.1: 2.Mean_Motion2 motion trace test 1

# 2.Mean_Motion (test2)

| Trace | 2.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | -0.0 |
| Std (mm) | 1.0 |
| 1st percentile (mm) | -2.5 |
| 99th percentile (mm) | 2.4 |
| Trace | 2.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.02 |

Figure 2.2: 2.Mean_Motion2 motion trace test 2

# 2.Mean_Motion2 (test3)

| Trace | 2.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | 0.0 |
| Std (mm) | 0.9 |
| 1st percentile (mm) | -2.3 |
| 99th percentile (mm) | 2.2 |
| Trace | 2.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.5 |
| 99th percentile (mm) | 1.5 |
| Latency (s) | 0.09 |

Figure 2.3: 2.Mean_Motion2 motion trace test 3

# 3.Mean_Motion_1 (test1)

| Trace | 3.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | -0.0 |
| Std (mm) | 1.1 |
| 1st percentile (mm) | -2.6 |
| 99th percentile (mm) | 2.8 |
| Trace | 3.Mean_Motion_6D |
| Mean (mm) | -0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.07 |

Figure 3.1: 3.Mean_Motion_1 motion trace test 1

# 3.Mean_Motion_1 (test2)

| Trace | 3.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | 0.0 |
| Std (mm) | 1.1 |
| 1st percentile (mm) | -2.4 |
| 99th percentile (mm) | 2.9 |
| Trace | 3.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.13 |

Figure 3.2: 3.Mean_Motion_1 motion trace test 2

# 3.Mean_Motion_1 (test3)

| Trace | 3.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | 0.0 |
| Std (mm) | 1.1 |
| 1st percentile (mm) | -2.4 |
| 99th percentile (mm) | 2.8 |
| Trace | 3.Mean_Motion_6D |
| Mean (mm) | -0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.12 |

Figure 3.3: 3.Mean_Motion_1 motion trace test 3

# 4.Mean_Motion_2 (test1)

| Trace | 4.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | -0.0 |
| Std (mm) | 1.0 |
| 1st percentile (mm) | -2.2 |
| 99th percentile (mm) | 2.7 |
| Trace | 4.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.09 |

Figure 4.1: 4.Mean_Motion_2 motion trace test 1

# 4.Mean_Motion_2 (test2)

| Trace | 4.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | -0.0 |
| Std (mm) | 1.0 |
| 1st percentile (mm) | -2.3 |
| 99th percentile (mm) | 2.5 |
| Trace | 4.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.07 |

Figure 4.2: 4.Mean_Motion_2 motion trace test 2

# 4.Mean_Motion_2 (test3)

| Trace | 4.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | 0.0 |
| Std (mm) | 1.2 |
| 1st percentile (mm) | -2.6 |
| 99th percentile (mm) | 2.6 |
| Trace | 4.Mean_Motion_6D |
| Mean (mm) | 0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.5 |
| 99th percentile (mm) | 1.4 |
| Latency (s) | 0.09 |

Figure 4.3: 4.Mean_Motion_2 motion trace test 3

# 1.Mean_Motion1 (test 1100s)

| Trace | 1.Mean_Motion_1D |
| --- | --- |
| Mean (mm) | -0.0 |
| Std (mm) | 0.7 |
| 1st percentile (mm) | -1.7 |
| 99th percentile (mm) | 1.6 |
| Trace | 1.Mean_Motion_6D |
| Mean (mm) | -0.0 |
| Std (mm) | 0.6 |
| 1st percentile (mm) | -1.4 |
| 99th percentile (mm) | 1.5 |
| Latency (s) | 0.28 |

Figure 5: 1.Mean_Motion1 motion trace test 1100s